The VIMS Wavelength and Radiometric Calibration White Paper

Roger N. Clark1, Robert H. Brown2, Dyer M. Lytle2

1 Planetary Science Institute, Lakewood, CO, 80227

2 Department of Planetary Sciences, University of Arizona, Tucson, AZ, 85721

Submitted to the NASA Planetary Data System

9/21/2016

Abstract

We derive a new radiometric calibration for the Visual and Infrared Mapping Spectrometer (VIMS) on Cassini. Since entering orbit in 2004, the VIMS instrument has undergone small shifts in the wavelength calibration of the spectrometer. The maximum shift is currently 10.4 nm (0.65 channels) and is estimated to be 0.71 channels at the end of mission. The wavelength shifts now require a time-dependent radiometric calibration technique to be deployed to preserve radiometric accuracy. Herein we quantify the time-dependent wavelength shift, and describe a compensatory scheme that will provide an accurate calibration for both specific intensity and I/F for VIMS measurements made during the Cassini Mission.

Introduction

The radiometric responsivity of a spectrometer or imaging instrument is complex, depending upon many factors, some of the most important being the aperture collection area, spectral transmission of the optical system, and the response function of the detector(s). The VIMS is an imaging spectrometer (Miller et al., 1996, Brown et al., 2004) whose spectral responsivity varies with angular position in the field of view and with wavelength. The IR chanel primary optical design incorporates Ritchey Cretien foreoptics, a reverse-Dahl-Kirkham collimator with a large central obscuration, a triply blazed reflection grating, Cassegrain camera optics and a linear, 256-element array of InSb detectors overlain with 4 order-sorting/thermal-blocking filter segments.  The VIS channel uses an off-axis Shafer telescope with an Offner design spectrometer. The flat-field response of the instrument varies with wavelength, and a flat-field image cube with 352 wavelengths is thus employed to correct the system response over the field of view relative to the instrument’s boresight pixel. The spectral bandpass of the instrument also varies slightly across the full range of VIMS’ spectral response, adding to the complexity of the calibration. Additional issues arise due to ringing in the instrument response function driven primarily by the blocking/order-sorting filters on the VIMS focal plane (in some cases ±10% or more [Figure 1]). Because of the ongoing wavelength shifts in the instrument, relative to when the instrument response function was measured on the ground and after the Cassini Jupiter flyby in late 2000, artifacts in radiometrically calibrated spectra result because the response function at any later time is actually sampled at different wavelengths than the wavelength set used in the ground measurements. The resulting artifacts are primarily caused by the use of a response function in multiplicative or divisive operations sampled using the original, ground wavelength set, and then used to calibrate to specific intensity or I/F instrument data obtained at a later time and consequently done with a different set of wavelengths. Below we describe our methodology to track the shifting wavelengths, and then apply a time-dependent, radiometric calibration that compensates for all the known effects.

Figure 1.  VIMS order-sorting/blocking filter transmission for the 1.6-3.0 micron segment.  The ringing in the transmission propagates into the radiometric calibration in a complex way because shift in wavelength sample different places in the response function.  (From Al-Jumaily, 1991).

VIMS Wavelengths Versus Time

The VIMS calibration pipeline and PDS-delivered data used a wavelength set and FWHM set measured during thermal vacuum testing at JPL before launch (Brown et al., 2004), but shifted by 1.3 channels as observed in Galilean Satellite data obtained during the Jupiter flyby (McCord et al., 2001). The calibration also incorporated a small, additional set of corrections to a few channels near the wavelength of the 4.25 micron CO2 absorption (Cruikshank et al., 2010). Here we employ terminology to describe the various corrections to the VIMS Radiometric Calibrations by using a numbering scheme RCnn (Radiometric Calibration nn). RC17 was the primary calibration used until June, 2016 (Clark et al., 2012), and has been used in all data processed through the VIMS calibration pipeline, and all files delivered to the PDS in the form of calibration files prior to the calibration described herein.

The wavelength calibration shift of the instrument since Cassini’s Saturn Orbit Insertion (SOI) is the primary driver of the need to employ a time-dependent radiometric calibration, and has been intensively studied since about 2013. The reader is directed to a separate white paper in the PDS describing the details of the VIMS time-dependent wavelength calibration, which we briefly describe here. Two methods for checking the VIMS wavelengths have been employed: (1) monitoring, as a function of time, the reflectivity in 3 windows of the Titan spectrum (1.6, 2.0, 2.8 microns; see the PDS white paper for details) by fitting Gaussian profiles to the window peaks; and (2) using the VIMS internal, wavelength-calibration, laser diode (central wavelength ~0.979 microns), and employing a fit to the Gaussian intensity profile of the calibration diode convolved with the response of 8 VIMS channels centered near peak intensity of the calibration diode. The result is the derivation of a time-dependent shift of the current wavelengths relative to the set extant at the time of Cassini’s SOI. Figures 2a, 2b show the results from both methods.

Figure 2a. VIMS wavelength shift versus time for two methods: Titan atmospheric windows (red points), and VIMS calibration diode (dark blue points).

Figure 2b. VIMS wavelength shift as measured by the calibration diode. The central, solid line is a linear regression, and the two dotted lines are ± 1 sigma as measured relative to the linear fit. It is not certain if the outlier points indicate real short term variations, but there are some indications that the fluctuations may be caused by temperature cycling of the VIMS spectrometer optics (see PDS white paper).

Because the shifts are small, we specify the shift on a yearly basis as described in Table 1. The shifts are relative to the new “standard 2004.0” wavelengths given in Table 2.

Calibration Equations

The basic equation we employ for calibrated, specific intensity, I, is:

- I=RDN−DarkflatCB,

where RDN is the instrument raw data numbers per second, Dark is the instrument measured dark current and thermal background (a dark/thermal background measurement is made after every VIMS scan line), and flat is the VIMS flat-field response at the given wavelength and position within the field of view normalized to the response of the boresight pixel (i.e., the relative response of the boresight pixel in the flat-field cube is 1 at every wavelength).   The units of equation 1 are: energy/time/area/bandpass/steradian .

C is defined as the calibration-multiplier vector which includes system transmission, grating efficiency, detector response, and divided by the exposure time (see equation 4, below).

- ,

where

λ = wavelength,

Δλ = full width, half maximum, FWHM,

Ω = solid angle = 2.5x10-7 steradian,

A = aperture area  IR:  A = 96.1  cm2, = 0.00961  m2 , VIS: A = 15.88 cm2, A = 0.001588 m2,

c = speed of light = 2.998x1010 cm/s = 2.998x108 m/s, and

h = Planck constant = 6.626x10-27 erg-s  = 6.626x10-34 J-s (= W s2).

The calibration to apparent reflectance, I/F, referred to that of a Lambert disk (a perfect diffuse reflector) is (note that the factor of π comes from the result that I/F for a Lambert disk is π-1):

- IF=πRDN−Darktflat×CBS,

where t is the exposure time – 0.004 second for the IR, or just the exposure time in seconds for the VIS channel, and S is the solar spectrum from Thompson et al. (2015), corrected to the flux extant at the heliocentric distance of the object in question.  The 4 millisecond IR exposure time correction is due to settling time after each mirror movement of the scanning secondary and was derived by McCord et al., (2004).

In the VIMS instrument, A and Ω are more complex (see Brown et al., 2005).  In standard resolution mode, the VIS channel bins 3x3 pixels to make a 0.5 x 0.5 mradian IFOV, and the IR IFOV is actually 0.5 x 0.25 mradian which gets moved 0.25 mradian half way through the integration to make a  0.5 x 0.5 mradian square pixel.  In high resolution mode, the VIS channel is 0.17 x 0.17 and the IR is 0.5 x 0.25 mradian.  The full telescope aperture, a 22.9 cm diameter Ritchey Cretien telescope, 800-mm focal length, has obscuration in both the Ritchey-Cretien secondary, and in the inverse Dahl-Kirkham collimator and Cassegrain spectrometer optics.  We use different constants to scale the relative multiplier (equation 4, below) to compensate for these factors to derive the value of C for each VIS and IR channel and each wavelength.

Derived Spectral Properties

The wavelength calibration and FWHM for VIMS were measured on the ground in a thermal vacuum chamber before launch (Brown et al., 2004). At the Jupiter fly-by it was determined that the IR wavelengths had shifted 1.3 channels from the ground calibration (McCord et al., 2001). That 1.3-channel shift from the ground calibration defined the VIMS standard wavelengths for post launch to RC17. No apparent shift occurred after the Jupiter fly-by and before Saturn orbit insertion.

Using the measured wavelength shifts from Table 1 we modeled the effect of the shift on solar port data from 2005 to 2012 (Figure 3). The model does not include order-sorting/blocking-filter response. Unfortunately, as the data in Figure 1 show, the analog plot of filter transmission made over 20 years ago is not precise enough for tracking the effects of the wavelength shifts.

Figure 3:  Solar port data from 2005 divided by data from 2012 (blue line) shows larger than predicted changes than just due to solar spectral structure.

Investigation of the causes of the discrepancy shown in Figure 3 reveal that, besides ringing in the order sorting filter transmission, fine structure in the original measured wavelengths of each vims channel and the derived FWHM display periodic structure which we traced to a periodic error in the grating position of the test monochromater used in the ground-based, thermal-vacuum testing of VIMS.  Figures 5 and 6 show the trends.

Figure 5. VIMS IR wavelengths divided by a linear fit to the end channels of the IR spectrometer (blue line). A cubic-spline fit (black line) is our new derivation: the new RC19 2004.0 “standard” wavelength set (values in Table 2). The blue line is a plot of the wavelengths derived in thermal vacuum testing before launch and shifted 1.3 channels (derived that the Jupiter fly-by) with a correction for the wavelengths around the 4.25 micron CO2 atmospheric absorption (Cruikshank et al. 2010).

Figure 6. VIMS IR-channel bandpasses (FWHM) in microns (blue line) derived during thermal vacuum testing before launch. A cubic-spline fit (black line) defines the new RC19 FWHM set (values in Table 2).

Calibration Multiplier

The calibration multiplier vector, C, from equation 1 is the major factor that remains to be determined.  Ideally it could be derived from measurements of a known, standard reflector illuminated with a source of known intensity, such as a calibration target on the spacecraft illuminated by the Sun, but such a device was descoped from VIMS in early Cassini budget cuts, thus we are left with only more approximate methods. Standard stars (that is, stars whose spectral specific intensities are known with the required precision) could be used, but because they are sub-pixel to VIMS, and because the VIMS pixel response function is variable within a pixel, unresolved objects like stars cannot be used with sufficient accuracy. Therefore, the only practical avenue that remains involves the use of known spectra of Saturn’s icy satellites and rings, Saturn itself, and Titan. Using a cross correlation of the spectra of the various objects, we were able to ascertain in general which “bumps and wiggles” in the response function were instrument related, and which were real features in the spectra of the various objects.

Furthermore, because the reflection spectra of solar-systems objects can be strong functions of their illumination geometry, the spectra of some objects were analyzed as a function of phase angle. For example, high-phase observations of Saturn's D and F-rings are expected to be dominated by diffraction, with muted spectral signatures of water ice, except near the water-ice vibrational fundamentals, where the real index of refraction becomes close to 1.0.  The derived, relative-multiplier vector for 2004.0 is shown in Figure 7.

VIMS has order-sorting filter gaps near 1.65, 3, and 3.9 microns (the upward spike in Figure 7), and they do not shift with time because they are fixed with respect to the detector array, whereas the spectrum of an object projected onto the detector array changes with time, and thus all other structure in spectral measurements does change in the IR channel of VIMS. The sampled wavelengths of the visible channel do not shift, thus its ground calibration is sufficient for calibration of the visible-channel data. The VIMS-IR multiplier vector (excluding the filter gap transmission) from Figure 7 was cubic-spline interpolated to the wavelength set derived for each year and given in Table 1. We then applied the effects of the filter-gap transmission.

Figure 7. The relative calibration multiplier vector, C, used in Equation 1. The vector is resampled with time as described in the text. To put it on the absolute level to be used as the variable C in equation 1, the vector must to be multiplied by the constant 29554.  See equation 4.

The results of the new calibration are illustrated in Figure 8. The RC17 calibration, which does not take into account the effects of shifting wavelengths shows increasing ringing (compare The D-ring 2006 RC 17 with F-ring 2015 RC17 spectra). The new RC19 calibration, which tracks the effects of the shifting wavelength sets, eliminates the ringing problems.

Remaining issues and uncertainties include:

1) The small rise near 2.75 microns may or may not be real.

2) The feature near 3 microns is on the edge of the order sorting filter gap and may or may not be real. If real the feature is most likely due to an N-H stretch (e.g. ammonia or ammonium compound).

3) The filter gap transmission around 1.65 and 3.9 microns may need adjustment.

Figure 8.  Example VIMS spectra using RC17 which was constant with time, and the new RC19 method described here.

Application/implementation of the time-dependent transfer functions to calibration of data

Calibration of VIMS uses equation 3, above.  The value of B from equation 2 is computed by selecting the wavelengths and FWHM for the time of the observation, and MKS units as follows.

Ω = 2.5x10-7 steradian.

c =  2.998x108 m/s.

A =  0.00961  m2 (VIS), 0.001588 m2 (IR).

h = 6.626x10-34 J-s (= W s2).

λ = wavelength in meters (ascii listings are in microns. So divide by 1000.0).

Δλ = FWHM in microns.

ASCII listing of the value of B for each VIMS channel is given in the Appendix.

In equation 3,

(4a)  CVIS =  29554. * calibration multiplier set from the appendix,

(4b)  CIR  =  107668. * calibration multiplier set from the appendix,

the calibration multiplier set is one of the files with a name of

vims.calibration.multiplier.RC19-20NN.N.txt

where 20NN.N is the year.    The constants in 4a and 4b apply to both standard and high resolution modes.

The Thompson et al. (2015) solar spectra convolved to VIMS are also given in the Appendix for each wavelength and calibration multiplier set.

The resulting units from equation 3 are dimensionless (I/F).  To get to radiance, multiply by the solar spectrum.  The supplied VIMS-convolved solar spectra are in units of Watts / m2 micron.

Plans for 2017-EOM

Wavelength calibration measurements using the instrument’s calibration diode will continue to be made every few days until the end of the Cassini mission. When all of the data regarding the shifts are collected, we will then derive a final set of wavelengths which will cover each year during the Cassini Orbital Tour. From those wavelengths will be derived the time-dependent calibration vectors to be applied to all of the VIMS raw data gathered during the Cassini Orbital tour. Those results will be deposited in the PDS 3-6 months after the end of the Cassini Mission. No further analysis will be carried out by the VIMS Team after the end of the Cassini Mission.

Appendices

A tar file, vims-calibration+wavelengths-vs-time-files.tar, contains ascii listings of the wavelengths,  calibration multiplier, the computed B values (equation 2), and the VIMS-convolved solar spectrum as a function of time.  The time increment is 1 year centered mid year.  The VIMS wavelength shift is small enough that this granularity is sufficient for most applications, as the shift between these calibrations is less than 0.0 channel.

Conclusions

The VIMS wavelengths are shifting small amounts per year (less than 0.1 channel), necessitating the need for a time-dependent calibration.  We have derived a new methodology that tracks this change and produces significantly better spectra with lower noise and lower artifacts.  This new calibration should enable greater confidence in observed spectral features, leading to detection of lower abundance components and therefore new and better science.

Table 1

Time    Shift       Shift       Add to standard wavelength set:

(nm)   Channels  (microns)

2004.0     0.0      0.000      0.0000 (= standard wavelength set)

2005.5     0.4      0.022      0.0004

2006.5     1.9      0.114      0.0019

2007.5     3.4      0.206      0.0034

2008.5     4.9      0.299      0.0049

2009.5     6.0      0.369      0.0060

2010.5     6.0      0.369      0.0060

2011.5     7.0      0.431      0.0070

2012.5     7.7      0.474      0.0077

2013.5     8.4      0.517      0.0084

2014.5     9.0      0.557      0.0090

2015.5     9.7      0.606      0.0097

2016.5    10.4      0.650     0.0104

2017.5    11.1      0.694     0.0111  (projected)

Table 2

VIMS Standard Wavelengths (2004.0)

Standard

wavelength       FWHM

(microns)      (microns)

1	       0.350540       0.007368

2	       0.358950       0.007368

3	       0.366290       0.007368

4	       0.373220       0.007368

5	       0.379490       0.007368

6	       0.387900       0.007368

7	       0.395180       0.007368

8	       0.402520       0.007368

9	       0.409550       0.007368

10	       0.417310       0.007368

11	       0.424360       0.007368

12	       0.431840       0.007368

13	       0.439190       0.007368

14	       0.446520       0.007368

15	       0.453720       0.007368

16	       0.461630       0.007368

17	       0.468410       0.007368

18	       0.476220       0.007368

19	       0.486290       0.007368

20	       0.489670       0.007368

21	       0.497770       0.007368

22	       0.506280       0.007368

23	       0.512220       0.007368

24	       0.519630       0.007368

25	       0.527660       0.007368

26	       0.534160       0.007368

27	       0.541560       0.007368

28	       0.549540       0.007368

29	       0.556140       0.007368

30	       0.563530       0.007368

31	       0.571310       0.007368

32	       0.578100       0.007368

33	       0.585480       0.007368

34	       0.593120       0.007368

35	       0.599380       0.007368

36	       0.607570       0.007368

37	       0.615050       0.007368

38	       0.622070       0.007368

39	       0.629400       0.007368

40	       0.637040       0.007368

41	       0.644080       0.007368

42	       0.651420       0.007368

43	       0.659100       0.007368

44	       0.666090       0.007368

45	       0.673420       0.007368

46	       0.681020       0.007368

47	       0.688030       0.007368

48	       0.695350       0.007368

49	       0.702880       0.007368

50	       0.710000       0.007368

51	       0.717330       0.007368

52	       0.724840       0.007368

53	       0.731980       0.007368

54	       0.739300       0.007368

55	       0.746760       0.007368

56	       0.753960       0.007368

57	       0.761280       0.007368

58	       0.768740       0.007368

59	       0.775950       0.007368

60	       0.783280       0.007368

61	       0.790720       0.007368

62	       0.797930       0.007368

63	       0.805220       0.007368

64	       0.812620       0.007368

65	       0.819890       0.007368

66	       0.827210       0.007368

67	       0.834630       0.007368

68	       0.841900       0.007368

69	       0.849220       0.007368

70	       0.856630       0.007368

71	       0.863910       0.007368

72	       0.871220       0.007368

73	       0.878630       0.007368

74	       0.885890       0.007368

75	       0.893860       0.007368

76	       0.900320       0.007368

77	       0.907870       0.007368

78	       0.915180       0.007368

79	       0.922540       0.007368

80	       0.929830       0.007368

81	       0.937130       0.007368

82	       0.944450       0.007368

83	       0.951770       0.007368

84	       0.959070       0.007368

85	       0.966380       0.007368

86	       0.973820       0.007368

87	       0.981000       0.007368

88	       0.988830       0.007368

89	       0.995880       0.007368

90	       1.002950       0.007368

91	       1.010050       0.007368

92	       1.016950       0.007368

93	       1.024710       0.007368

94	       1.031950       0.007368

95	       1.038650       0.007368

96	       1.045980       0.012480   End of VIS channel

97	       0.884210       0.012878

98	       0.900753       0.012767

99	       0.916924       0.012507

100	       0.933078       0.013169

101	       0.949803       0.012869

102	       0.965683       0.012728

103	       0.982262       0.013370

104	       0.998820       0.012790

105	       1.014790       0.012748

106	       1.031320       0.013186

107	       1.047550       0.012847

108	       1.065410       0.013136

109	       1.081830       0.013063

110	       1.098060       0.012686

111	       1.113960       0.012828

112	       1.130240       0.013111

113	       1.146950       0.013322

114	       1.163700       0.013266

115	       1.179960       0.012968

116	       1.196220       0.013018

117	       1.212460       0.012921

118	       1.228590       0.013031

119	       1.244920       0.013401

120	       1.261660       0.013631

121	       1.278130       0.013372

122	       1.294820       0.013121

123	       1.310910       0.013101

124	       1.326950       0.013146

125	       1.343240       0.013389

126	       1.359520       0.013663

127	       1.376950       0.013366

128	       1.393260       0.012821

129	       1.409400       0.013147

130	       1.425570       0.013137

131	       1.441840       0.013216

132	       1.458410       0.013480

133	       1.475140       0.013610

134	       1.491690       0.013066

135	       1.507940       0.013063

136	       1.524210       0.012992

137	       1.540350       0.013059

138	       1.556740       0.013388

139	       1.573610       0.014011

140	       1.590180       0.013901

141	       1.602280       0.008457   order-sorting filter change

142	       1.625230       0.032000   order-sorting filter change

143	       1.641600       0.009862   order-sorting filter change

144	       1.655670       0.013304

145	       1.672380       0.013532

146	       1.689010       0.013253

147	       1.705360       0.013300

148	       1.721750       0.013068

149	       1.738020       0.013084

150	       1.754360       0.013155

151	       1.771050       0.013455

152	       1.787710       0.013080

153	       1.804010       0.013090

154	       1.820040       0.012902

155	       1.836160       0.012985

156	       1.852880       0.013531

157	       1.869330       0.012939

158	       1.886790       0.012600

159	       1.902610       0.013058

160	       1.919160       0.013059

161	       1.935450       0.013127

162	       1.951910       0.013498

163	       1.968710       0.013615

164	       1.985310       0.013293

165	       2.001670       0.013209

166	       2.017810       0.013294

167	       2.034240       0.013415

168	       2.050910       0.013889

169	       2.067570       0.013472

170	       2.084000       0.013579

171	       2.100340       0.013428

172	       2.116670       0.013719

173	       2.133370       0.013943

174	       2.150180       0.013787

175	       2.166520       0.013547

176	       2.182880       0.013600

177	       2.199200       0.013571

178	       2.215910       0.014009

179	       2.232820       0.013918

180	       2.249520       0.013700

181	       2.266220       0.013600

182	       2.282380       0.014012

183	       2.299210       0.013974

184	       2.316120       0.014211

185	       2.333250       0.014287

186	       2.350430       0.014407

187	       2.367650       0.014286

188	       2.384720       0.014294

189	       2.401560       0.014079

190	       2.418200       0.013921

191	       2.434710       0.013829

192	       2.450970       0.013748

193	       2.467230       0.013784

194	       2.483600       0.014044

195	       2.500020       0.014293

196	       2.516590       0.014306

197	       2.532920       0.013704

198	       2.549160       0.013918

199	       2.564370       0.011963

200	       2.581760       0.013610

201	       2.598070       0.014726

202	       2.615080       0.012722

203	       2.630000       0.011283

204	       2.646500       0.013711

205	       2.661460       0.012674

206	       2.680850       0.016119

207	       2.696200       0.014697

208	       2.712050       0.015964

209	       2.732700       0.012786

210	       2.747700       0.018701

211	       2.763050       0.016296

212	       2.781180       0.013689

213	       2.798890       0.014400

214	       2.816060       0.015083

215	       2.832470       0.014680

216	       2.849540       0.014842

217	       2.866090       0.015667

218	       2.882420       0.015534

219	       2.898780       0.015325

220	       2.915400       0.015088

221	       2.931430       0.015720

222	       2.947260       0.015350

223	       2.963270       0.015716   order-sorting filter change

224	       2.977200       0.015512   order-sorting filter change (~30% transmission)

225	       3.000720       0.012919   order-sorting filter change

226	       3.013820       0.015570

227	       3.029700       0.015398

228	       3.048060       0.014922

229	       3.064460       0.015466

230	       3.080360       0.015882

231	       3.096890       0.015851

232	       3.112130       0.015753

233	       3.129620       0.016580

234	       3.146670       0.015851

235	       3.163040       0.016127

236	       3.179740       0.016115

237	       3.197080       0.015685

238	       3.213640       0.015830

239	       3.231500       0.016740

240	       3.248060       0.017771

241	       3.265610       0.016161

242	       3.282980       0.016285

243	       3.299460       0.016286

244	       3.316190       0.015816

245	       3.333380       0.015203

246	       3.349810       0.016500

247	       3.365640       0.015590

248	       3.381830       0.015717

249	       3.398720       0.016471

250	       3.415460       0.016457

251	       3.431780       0.016343

252	       3.448740       0.015852

253	       3.464750       0.015634

254	       3.481370       0.015608

255	       3.497950       0.015779

256	       3.512840       0.016141

257	       3.530150       0.015057

258	       3.546640       0.016643

259	       3.562740       0.016735

260	       3.580340       0.016474

261	       3.596100       0.017033

262	       3.613870       0.020159

263	       3.630850       0.018293

264	       3.648530       0.017622

265	       3.665220       0.018895

266	       3.682830       0.018505

267	       3.699530       0.019496

268	       3.717430       0.018635

269	       3.734390       0.019045

270	       3.751030       0.019296

271	       3.767630       0.017966

272	       3.784440       0.019006

273	       3.800830       0.018599

274	       3.817420       0.018210

275	       3.834720       0.019856

276	       3.851410       0.018125

277	       3.861840       0.015574   order-sorting filter change

278	       3.881670       0.023959   order-sorting filter change

279	       3.898590       0.020270

280	       3.914780       0.021217

281	       3.930690       0.020631

282	       3.947620       0.019721

283	       3.963750       0.020799

284	       3.980150       0.021142

285	       3.996720       0.021846

286	       4.012800       0.021142

287	       4.029440       0.021531

288	       4.047300       0.021598

289	       4.062950       0.022566

290	       4.080861       0.021479

291	       4.097430       0.022433

292	       4.114500       0.022013

293	       4.131830       0.022290

294	       4.148830       0.022294

295	       4.166440       0.020424

296	       4.183199       0.021180

297	       4.200100       0.019057

298	       4.217000       0.017383

299	       4.233700       0.022866

300	       4.250500       0.021600

301	       4.267300       0.021600

302	       4.284000       0.021600

303	       4.300600       0.021600

304	       4.317101       0.021600

305	       4.333600       0.021600

306	       4.350200       0.019916

307	       4.366500       0.021335

308	       4.382900       0.020260

309	       4.397930       0.019563

310	       4.415370       0.021034

311	       4.431720       0.019802

312	       4.447720       0.019867

313	       4.465730       0.019735

314	       4.482400       0.019931

315	       4.499511       0.021189

316	       4.515910       0.020529

317	       4.533790       0.019303

318	       4.551870       0.019921

319	       4.567970       0.020376

320	       4.585560       0.020113

321	       4.602900       0.019105

322	       4.620100       0.020267

323	       4.636150       0.017409

324	       4.654160       0.019612

325	       4.670340       0.018281

326	       4.687210       0.019814

327	       4.702900       0.017902

328	       4.719561       0.020080

329	       4.737060       0.018831

330	       4.753510       0.018017

331	       4.770310       0.015955   hot pixel

332	       4.786730       0.018821

333	       4.803490       0.017274

334	       4.819521       0.018179

335	       4.835771       0.018545

336	       4.852920       0.018106

337	       4.869401       0.018799

338	       4.885530       0.019556

339	       4.902650       0.018114

340	       4.919831       0.018570

341	       4.936851       0.017740   slightly noisier (1.5x  at low signal)

342	       4.953890       0.018779

343	       4.971780       0.018266

344	       4.988960       0.020001

345	       5.005760       0.018402

346	       5.022400       0.018621

347	       5.040781       0.016783

348	       5.057340       0.019510

349	       5.074020       0.017953

350	       5.091060       0.020883  slightly noisier (2x at low signal)

351	       5.106800       0.015704

352	       5.122500       0.016

References

Al-Jumaily, G.A., E. Miller, and N. Raof, 1991, C/C VIMS Filter requirements, JPL VIMS Design File Memorandum, 6-25-1991.

Brown, R.H., Baines, K.H., Bellucci, J.P., Bibring, B.J., Buratti, E., Bussoletti, F., Capaccioni, P., Cerroni, P., Clark, R.N., Coradini, D.P., Cruikshank, P., Drossart, Y., Formisano, R., Jaumann, Y., Langevin, D.J., Matson, T.R., 2005, The Cassini visual and infrared mapping spectrometer investigation: Space Science Reviews, 115 (1-4), 111-168.

Clark, R. N., D. P. Cruikshank, R. Jaumann, R. H. Brown, K. Stephan, C. M. Dalle Ore, K. E. Livo, N. Pearson, J. M. Curchin, T. M. Hoefen, B. J. Buratti, G. Filacchione, K. H. Baines, and P. D. Nicholson, 2012, The Composition of Iapetus: Mapping Results from Cassini VIMS, Icarus, 218, 831-860.

Cruikshank, D. P., A. W. Meyerb and, R. H. Brownc and, R. N. Clark, R. Jaumann, K. Stephan, C. A. Hibbitts, S. A. Sandford, R. M.E. Mastrapa, G. Filacchione, C. M. Dalle Ore, P. D. Nicholson, B. J. Buratti, T. B. McCord, R. M. Nelson, J. B. Dalton, K. H.  B., and D. L. Matson, 2010, Carbon dioxide on the satellites of Saturn: Results from the Cassini VIMS investigation and revisions to the VIMS wavelength scale, Icarus, 206, 561–572, doi:10.1016/j.icarus.2009.07.012 .

McCord, T. B., Brown, R., Baines, K., Bellucci, G., Bibring, J.-P., Buratti, B., Capaccioni, F., Cerroni, P., Clark, R.N., Coradini, A., Cruikshank, D., Drossar, P., Formisano, V., Jaumann, R., Langevin, Y., Matson, D., Mennella, V., Nelson, R., Nicholson, P., Sicardy, B., Sotin, C., Hansen, G., Hibbitts, C., 2004, Cassini VIMS observations of the Galilean satellites, including the VIMS calibration procedure: Icarus, Vol. 172, p. 104-126.

Miller, Edward A., Gail Klein, David W. Juergens, Kenneth Mehaffey, Jeffrey M. Oseas, Ramon A.

Garcia, Anthony Giandomenico, Robert E. Irigoyen, Roger Hickok, David Rosing, Harold R. Sobel,

Carl F. Bruce, Jr., Enrico Flamini, Romeo DeVidi, Francis M. Reininger, Michele Dami, Alain

Soufflot, Yves Langevin, Gerard Huntzinger, 1996, The Visual and Infrared Mapping Spectrometer for Cassini , Proc. SPIE  2803, Cassini/Huygens: A Mission to the Saturnian Systems, 206-220 (October 7, 1996), doi: 10.1117/12.253421 .

Thompson, D. R., F. C. Seidel, B. C. Gao, M. M. Gierach, R. O. Green, R. M. Kudela, and P. Mouroulis , 2015, Optimizing irradiance estimates for coastal and inland water imaging spectroscopy, Geophys. Res. Lett., 42, 4116–4123, doi:10.1002/2015GL063287.